**Supplemental Table 2:** Programmed trigger of the BioMonitor 2-AF for arrhythmic events.

|  |  |
| --- | --- |
| **Atrial Fibrillation (AF)** | Irregular RR intervals and absence of P waves (confirmation interval for detection: 2 min) |
| **Asystole** | Absence of electrical activity (no P wave, QRS complex) for at least 5 sec |
| **Bradycardia** | Heart rate below 40 bpm (RR < 1500 msec) for at least 15 sec |
| **High Ventricular Rate (HVR)** | Ventricular rate of at least 180 bpm (confirmation interval for detection: 24 RR-interval, RR < 333 msec) |